## Slide 1
Does the Use of Ceramic Femoral Head Versus Metal Femoral Head Improve the Outcome of Primary Total Hip Arthroplasty?
Terry Clyburn, MD
Professor Orthopedic UT Houston (ret)
Total Joint Surgeon, Houston Methodist Hospital

## Slide 2
Experts
Dragan Radoicic, MD PhD
Yihe Hu, MD PhD
Terry Clyburn, MD
Koos Jordaan, MD
Jasmin Ciriviri, MD
Eoin Sheehan, MD
Liaisons
Elizabeth Abe, BS
Graham Goh, MD

## Slide 3
Metal head on cross-linked polyethylene (M-on-XLP) has the longest follow-up and excellent outcomes, even in patients under 50 y/o.
Ceramic heads have improved wettability and surface smoothness and are thought to be more resistant to third body wear.
There is no current consensus on the best material for the femoral head.
Why is this topic important?

## Slide 4
55 original studies reviewed (2001-2022)
-Including 4,063 patients
3 Systemic Reviews With Meta-Analysis
2001-2022
Literature Review
Mesh Terms
Total hip arthroplasty
Hip prosthesis
Ceramic
Metal
Polyethylene
Arthroplasty, replacement hip
Prosthesis design
Prosthesis failure
Reoperation total hip
Risk Factors total hip
Follow-up studies total hip

## Slide 5
THA is highly successful, cost-effective for end-stage OA
Limitations to long-term implant survivorship1-3
  Polyethylene wear
  Osteolysis
  Adverse local tissue reaction
Highly cross-linked polyethylene (XLPE) acetabular liners first introduced in 1990s4
  Superior resistance to wear5
  Largely replaced conventional polyethylene liners5
Background

## Slide 6
Ceramic femoral heads6-11
  Long-term survivorship
  Decreased risk of PJI
  Minimal risk of component fracture
Metal-on-polyethylene (MoP) versus ceramic-on-polyethylene (CoP)
  Cost + personal experience cited as factors influencing surgeon’s implant selection12
  Several early studies have shown no difference in wear rates11,13-14
Literature Findings

## Slide 7: Literature Findings
Mertz et.al. (2023)- “Ceramic Femoral Heads Exhibit Lower Wear Rates Compared to Cobalt Chrome: A Meta-Analysis” ; 4th generation ceramic vs. metal head on XLPE
  36 studies included in review and Four studies included in meta-analysis; (287 patients, 172 ceramic heads).
  Pooled weighted data from all 36 studies- CoCr- 0.063mm/year
  Ceramic- 0.047 mm/yr ( P<0.01)
  CoCr- 0.029 mm/yr greater wear than ceramic heads.

## Slide 8
Question
Does the Use of Ceramic Femoral Head Versus Metal Femoral Head Improve the Outcome of Primary Total Hip Arthroplasty?
Recommendation
Fourth-generation ceramic femoral heads with highly cross-linked polyethylene liner may be associated with decreased linear wear rates and improved patient-reported outcomes, when compared to other bearing surfaces.
Recommendation; Moderate Level